# OBR Fields

# Background

The use of additional fields in the OBR segment has been proposed.

*Beckman comments in blue

# Fields

- OBR-13 Relevant Clinical Information: Could/should NTE segments be used instead? NTE segments are now supported by LAW. Assumption is that this information contains comments and notes intended for humans, and will not be interpreted by a machine. Otherwise OBX segments are recommended for structured data. Reject, recommend use of NTE segments.

  - Beckman needs a machine readable diagnosis for processing of decision rules.  This field is described in HL7 documentation as diagnosis and is suitable for providing the required information.

  - Beckman recommends that OBR-13 field be added to LAW.

  - Riki Merrick: for structured data of clinical importance should use OBX after the OBR or the SPM they refer to – in v2.8.1 there is a new field OBX-29 to indicate, if OBX carries a result, response to AOE or other clinical information – that is what LRI will use next DST update.

  - IHE LAB recommendation

    - Pre-adopt OBX-29 to identify type of OBX segment

    - Use OBX segment for clinical information

    - OBSERVATION segment group of LAB-28 (OML^O33) can be used to carry clinical information

    - Expand discussion near line 3370 to expand discussion on potential uses of OBSERVATION segment group.

- OBR-16 Ordering Provider: Beckman uses the physician name fields. LAW only uses <ID Number>. Beckman and bioMerieux will check what is needed and suggest changes. Recommendation is to add Physician Name Field. This will be enhanced interface data.

  - Beckman recommends adding the physician name fields.

  - Riki Merrick: Consider adding Assigning authority to the ID, too (XCN.9).

  - IHE LAB recommendation

    - Expand OBR-16 to include physician name field

    - Include Assigning Authority as an optional field

- OBR-18 Placer Field 1: Beckman specific data. Beckman started with existing interface and found no place to map requisition field, so used this field. Received from LIS, so use it in reply. May no longer be needed with AWOS ID. Beckman to confirm it is still needed. Assumption is no longer required due to use of AWOS ID.

  - The requisition number can also be entered at the Analyzer and sent to the LIS.  Beckman will send the requisition number in a DxH defined OBX segment to the LIS.

  - Beckman recommends not adding OBR-18.

  - IHE LAB Recommendation

    - Leverage pre-adoption of OBX-29

    - Send OBX segments containing this, and other AOE type information, along with other AWOS results

    - Update discussions in LAW to highlight use of OBX segments in LAB-29 to carry information in addition to results. Need to discuss values for fields when OBX is not a “result” segment.

    - OBX segments with an OBX-29 set to something other than “result” can be discarded by the Analyzer Manager

- OBR-20 Filler Field 1: Beckman specific data. Manual edit of Specimen ID or tube position ID. Made recommendation to indicate original AWOS is an exception, and send result for unsolicited AWOS Status Update with actual Specimen or tube position ID. Beckman will review and consult with IHE LAB. Should not need. No NTE would be needed in the unsolicited result. Recommend not adding this field.

  - Beckman agrees with the above recommended actions when a manual edit of a Specimen ID or Tube Position ID is edited.

  - The manual edit flags will also be sent to the LIS in DxH defined OBX segments.

  - Beckman recommends not adding OBR-20.

  - IHE LAB recommendation

    - Send exception for original AWOS, and then send unsolicited AWOS for new container ID, tube position, etc.

    - Use OBX, with proper OBX-29 value, to document manual edit information.

- OBR-21 Filler Field 2: Beckman specific data. Quality Control information. Information needed to fully process control information. Maybe split into OBX and INV. Beckman will determine if INV could be used. Recommend using INV if possible and not adding this field.

  - The information from this field will be sent to the LIS in DxH defined OBX segments.

  - Beckman recommends not adding OBR-21.

  - IHE LAB recommendation

    - Use OBX, with proper OBX-29 value, to document additional QC information.

- OBR-22 Results Rpt/Status Chng – Date/Time: Release date/time. Could be the same as OBX-19. Beckman will review use of time stamps in OBR, ORC, or OBX to determine if this field is needed. Recommend not adding if existing fields are equivalent.

  - OBX-19 field is sufficient.

  - Beckman recommends not adding OBR-22.

  - Riki Merrick: This field will change with every update, while OBX-19 may not change, but between managers and analyzers that may not be the case – for sure needed from LIS to end user.

  - IHE Lab Recommendation

    - No changes necessary.

- OBR-25 Result Status: Release status. Beckman will review and determine if ORC-5 and OBX-11 suffice. Recommend not adding this field if information provided by ORC-5 and OBX-11.

  - OBX-11 field is sufficient.

  - Beckman recommends not adding OBR-25.

  - IHE Lab Recommendation

    - No changes necessary.

- OBR-32 Principal Result Interpreter: Operator ID of reviewer. Recommend reviewing use of OBX-16 and not adding this field.

  - OBX-16 field is sufficient.

  - Beckman recommends not adding OBR-32.

  - IHE Lab Recommendation

    - No changes necessary.

- OBR-46 Placer Supplemental Service Information: Number of slides created. Siemens recommended TCD extensions (Siemens extension TCD-10, or other new field in TCD) be used to specify number of runs. Slides have been numbered. TQ1- Quantity is a possibility. Need to add ability to send replicates. May be unique to an instrument, so evaluate this as well. Ultimately may need to propose a change to HL7 OO. Siemens has other extensions for TCD that should be reviewed.

  - DxH will use the TQ1-2 Quantity field to hold the number of slides.

  - Beckman recommends not adding OBR-46.

  - IHE LAB recommendation

    - Align with ordering replicates use case

    - Anticipate using single AWOS for each replicate, rather than TQ1-2