Minutes
Discrete Positioning and Delivery Workflow (DPDW)

Conference Call
December 16, 2014
11:00am – 12:30pm EST

DPDW Subgroup Chair:

Thomas Schwere, Varian Medical Systems
(thomas.schwere@varian.com)
IHERO Task Force Co-Chairs

Dick Fraass, Ph.D., FAAPM, FASTRO, FACR

John Buatti, MD
Mission Statement:  The American Society for Radiology Oncology (ASTRO) has formed a multi-society Task Force to undertake an initiative to promote the Integration of the Healthcare Enterprise (IHE) – Radiation Oncology (RO), fostering seamless connectivity and integration of radiotherapy equipment and the patient health information systems. The Task Force will include members from ASTRO, RSNA, American Association of Physicists in Medicine (AAPM), the American College of Radiology (ACR) and the Medical Imaging and Technology Alliance (MITA). In addition, members of the International community have also been invited to participate in IHE-RO. The IHE-RO Task Force, in close collaboration with radiotherapy product manufacturers, will develop appropriate integration profiles for radiation therapy and setup a demonstration of seamless communication among the full array of radiotherapy products.
In Attendance:
Crystal Carter (ASTRO)

Luiza Kowalczyk (Medical Imaging)
Ulrich Busch (Varian)

Thomas Schwere (Varian)
Sanjay Bari (Elekta)
Martin Vonach (Brainlab)

Harold Beunk (ICT)
1. Call to Order

The meeting was started at 11:00 EST.
2. Transactions
While incorporating the changes for RO-DPD-223, T. Schwere wondered what synchronization mechanism is used between RO-DPD-201 (Retrieve Device Position Information) and the actual acquisition of the positioning information (e.g. RT Images). The group discussed this:
· Currently no synchronization is foreseen.

· In any case the acquired data is tagged with the position information (e.g. table top position in the RT Image) allowing the calculation of an absolute target position.

· The description of RO-DPD-201 should explicitly say that the device position information should be retrieved together with the acquisition of the positioning data as simultaneous as possible. However, this does still not guarantee that the couch is moved in between.

· Couch locking options were discussed:

· RO-DPD-201 could implicitly lock the couch from any movement. This would require an additional transaction to unlock the couch again.

· However, manual couch movements in case of emergency must always be possible.
· RO-DPD-201 should be possible at any point in time and locking is not always desired.
· Locking could also happen on a real-time I/F (similar to the monitoring scenario).
· DPDW cannot change well-established workflows, e.g. off-center setup with a subsequent centering of the couch for a CBCT acquisition. Therefore locking of the couch seems to be problematic.

· Positioning data should be robust against couch movements.

· The group agreed on adding an open item for this to be discussed in the face-to-face meeting.

M. Vonach went through transaction RO-DPD-205:
· What UPS procedure code should be used for the RT Patient Position Correction? The group agreed on using 121722 "RT Patient Position Adjustment" from CID 9241 for that purpose.
· How to deal with procedures that are not supported by a particular performing device:

· Support of a procedure is mainly based on the instruction (the UPS in general might be supported).

· If a procedure/instruction is not supported the performing device shall set the UPS to CANCELED together with a reason for cancelation.

· Fig. 5.4-7: To reduce duplications (and probability of inconsistencies) consider using callouts to other diagrams.
3. General Design Topics
3.1. New Workflow Notification Service
Application Hosting Service:

H. Beunk is still working on the memo and will distribute it prior to the next meeting. The memo will be discussed in the next call.

4. Face-to-Face Meeting
There are five participants so far.
5. Other Business
5.1. Next DPDW TCon
· Tuesday, January 27, 2014
11:00am – 12:30pm EST.
6. Adjournment

The meeting was adjourned at 12:15 EST.

Appendix A: Administration and Process Information
Documents are published at the following locations. If you have problems in accessing the document, please contact the Chair (ulrich.busch@varian.com).
Process of Authoring:

Steps:

1. Download a local copy of the document from locations below
2. Open this copy and remove all change bars

3. Ensure, that Changes Bars are switched on

4. Make your changes
5. Provide the updated version to the Chair


Location of Documents:
DPDW Subgroup Minutes
http://wiki.ihe.net/index.php?title=RO_DPDW_WorkingGroup
DPDW Profile
The DPDW Profile is an IHE-RO document.

The current version is available in the IHE-RO Org Wiki:
http://www.ihe-ro.org/
Please find the current document under this page:

http://www.ihe-ro.org/doku.php?id=doc:profiles
Supp 160
DICOM Supplement 160 (Patient Positioning and Workflow) in s DICOM WG-07 document. 
The current version is available at the DICOM ftp server:

ftp://d9-workgrps:goimagego@medical.nema.org/MEDICAL/Private/Dicom/WORKGRPS/WG07/Supp160

Mailing List:
The mailing list for the DPDW subgroup is:
iherodpdw2014@mail.aapm.org
Appendix B: Task Assignments
Per end this TCon (2014-09-30).
	No
	TX / Area
	Old Number
	Title
	Group
	Owner

	1
	./.
	./.
	Use Case Delivery-Device Independent Imaging
	 
	David Wikler

	2
	RO-DPD-200
	RO-DPD-01
	Worklist Query for Positioning Acquisition
	Acquisition
	Martin Vonach

	3
	RO-DPD-201
	RO-DPD-02
	Retrieve Device Position Information
	Acquisition
	Martin Vonach

	4
	RO-DPD-202
	RO-DPD-03
	Request RT Patient Position Correction
	Correction
	Martin Vonach

	5
	RO-DPD-203
	RO-DPD-04
	Store RT Patient Position Modification Instruction
	Correction
	Martin Vonach

	6
	RO-DPD-204
	RO-DPD-05
	Store RT Repositioning Results to Object Storage
	Correction
	Martin Vonach

	7
	RO-DPD-205
	RO-DPD-06
	Worklist Query for Repositioning
	Correction
	Martin Vonach

	8
	RO-DPD-206
	RO-DPD-07
	Notify on Radiation Delivery Status Change
	Delivery
	Thomas Schwere, Sanjay Bari

	9
	RO-DPD-207
	RO-DPD-08
	Retrieve RT Patient Position Correction Instruction
	Correction
	Martin Vonach

	10
	RO-DPD-208
	RO-DPD-09
	Subscribe/Unsubscribe to Treat UPS Status
	UPS Notification
	Thomas Schwere

	11
	RO-DPD-209
	RO-DPD-10
	Notify on Radiation State
	Delivery
	Thomas Schwere, Sanjay Bari

	12
	RO-DPD-210
	RO-DPD-11
	Retrieve Positioning Acquisition Results
	Registration
	Chris Pauers

	13
	RO-DPD-211
	RO-DPD-12
	Worklist Query for Positioning Registration
	Registration
	Chris Pauers

	14
	RO-DPD-212
	RO-DPD-13
	Worklist Query for Position Monitoring
	Monitoring
	Stephen Phillips 

	15
	RO-DPD-213
	RO-DPD-16
	Store Monitoring Results to Object Storage
	Monitoring
	Stephen Phillips 

	16
	RO-DPD-214
	RO-DPD-17
	UPS Final Update at Session Termination
	Framework
	Thomas Schwere, Sanjay Bari

	17
	RO-DPD-215
	RO-DPD-18
	UPS Completed / Cancelled at Session Termination
	Framework
	Thomas Schwere, Sanjay Bari

	18
	RO-DPD-216
	RO-DPD-19
	Indicate Ready for Monitoring
	Monitoring
	Stephen Phillips 

	19
	RO-DPD-217
	RO-DPD-20
	Notify Device to start UPS
	UPS Notification
	Thomas Schwere, Sanjay Bari

	20
	RO-DPD-218
	RO-DPD-21
	Create Positioning Acquisition and Positioning Registration UPS
	Workflow
	Thomas Schwere

	21
	RO-DPD-219
	RO-DPD-22
	Create Treat UPS and Radiation Delivery Instruction for Continuation
	Workflow
	Thomas Schwere

	22
	RO-DPD-220
	RO-DPD-23
	Notify Treatment Session Actors on Starting Session
	UPS Notification
	Thomas Schwere, Sanjay Bari

	23
	RO-DPD-221
	RO-DPD-24
	Notify Device to stop UPS
	UPS Notification
	Thomas Schwere, Sanjay Bari

	24
	RO-DPD-222
	RO-DPD-25
	UPS Progress Update for Discrete non-Treatment Steps
	UPS Notification
	Thomas Schwere, Sanjay Bari

	25
	RO-DPD-223
	RO-DPD-26
	Worklist Query for Positioning Correction Reconciliation
	Registration
	Chris Pauers

	26
	RO-DPD-224
	RO-DPD-27
	External Verification
	External Verification
	Sanjay Bari

	27
	RO-DPD-225
	./.
	Notify Device to resume UPS 
	Monitoring
	Stephen Phillips 

	28
	RO-DPD-226
	./.
	Create new Positioning UPS
	Monitoring
	Stephen Phillips 

	29
	RO-DPD-227
	./.
	UPS Final Update after Positioning Information Acquisition
	Workflow
	

	30
	RO-DPD-228
	./.
	UPS Final Update after Treatment Interruption
	Workflow
	Thomas Schwere


