# IHE Change Proposal

Tracking information:

| IHE Domain | LAB |
| --- | --- |
| Change Proposal ID: | CP-LDA-NTE-in-RSP.docx |
| Change Proposal Status: |  |
| Date of last update: |  |
| Person assigned: |  |

Change Proposal Summary information:

| Add NTE segments to RSP^WOS |  |
| --- | --- |
| Submitter’s Name(s) and e-mail address(es): | Dmytro Rud, dmytro.rud@roche.com |
| Submission Date: | 08 May 2014 |
| Integration Profile(s) affected: | LDA |
| Actor(s) affected: | Automation Manager, Laboratory Device |
| IHE Technical Framework or Supplement modified: | LAB |
| Volume(s) and Section(s) affected: | Volume 2a |
| Rationale for Change: - The Automation Manager has no possibility to provide order comments in RSP^WOS messages. |  |

Section 3.7.5.2

Add marked rows to the table 3.7.5.5-2.

NB: Group SPECIMEN in RSP^WOS is a proprietary IHE LAB extension to HL7, so no HL7 change request is required.

Table ‎3.7.5.2-2: RSP^WOS^RSP_K11

| Segment | Meaning | Usage | Card. | HL7 chapter |
| --- | --- | --- | --- | --- |
| MSH | Message header | R | [1..1] | 2 |
| [{SFT}] | Software Segment | O | [0..*] | 2 |
| MSA | Message Acknowledgement | R | [1..1] | 2 |
| [ERR] | Error | O | [0..1] | 2 |
| QAK | Query Acknowledgement | R | [1..1] | 5 |
| QPD | Query Parameter Definition | R | [1..1] | 5 |
| { | --- SPECIMEN begin | C | [0..*] |  |
| SPM | Specimen | R | [1..1] | 7 |
| [{OBX}] | Observation related to specimen | O | [0..*] | 7 |
| [{SAC}] | Specimen Container | RE | [0..1] | 13 |
| [ | --- PATIENT begin | O | [0..1] |  |
| PID | Patient Identification | R | [1..1] | 3 |
| [{OBX}] | Observation related to the patient | O | [0..*] | 7 |
| ] | --- PATIENT end |  |  |  |
| { | --- ORDER begin | R | [1..1] |  |
| ORC | Common Order | R | [1..1] | 4 |
| [{ NTE }] | Notes and comments (for Common Order) | O | [0..*] | 2 |
| [{TQ1}] | Timing/Quantity | RE | [0..1] | 4 |
| [ | --- OBSERVATION REQUEST begin | O | [0..1] |  |
| OBR | Observation Request | R | [1..1] | 4 |
| [ TCD ] | Test Code Details | O | [0..1] | 13 |
| [{ NTE }] | Notes and comments (for Observation Request) | O | [0..*] | 2 |
| [{ | --- OBSERVATION begin | O | [0..*] |  |
| OBX | Observation/Result | R | [1..*] | 7 |
| [ TCD ] | Test Code Detail | O | [0..*] | 13 |
| [{ NTE }] | Notes and Comments (for Results) | O | [0..*] | 2 |
| }] | --- OBSERVATION end |  |  |  |
| ] | --- OBSERVATION REQUEST end |  |  |  |
| [{ | --- PRIOR RESULT begin | O | [0..*] |  |
| PV1 | Patient Visit – previous result | R | [1..1] | 3 |
| { | --- ORDER PRIOR begin | R | [1*] |  |
| ORC | Common order – previous result | R | [1..1] | 4 |
| OBR | Order detail – previous result | R | [1..1] | 4 |
| { | --- OBSERVATION PRIOR begin | R | [1..*] |  |
| OBX | Observation/Result – previous result | R | [1..1] |  |
| [{NTE}] | Comment of the result | C | [0..*] | 2 |
| } | --- OBSERVATION PRIOR end |  |  |  |
| } | --- ORDER PRIOR end |  |  |  |
| }] | --- PRIOR RESULT end |  |  |  |
| } | --- ORDER end |  |  |  |
| } | --- SPECIMEN end |  |  |  |