# OBR Fields

# Background

The use of additional fields in the OBR segment has been proposed.

# Fields

- OBR-13 Relevant Clinical Information: Could/should NTE segments be used instead? NTE segments are now supported by LAW. Assumption is that this information contains comments and notes intended for humans, and will not be interpreted by a machine. Otherwise OBX segments are recommended for structured data. Reject, recommend use of NTE segments.

- OBR-16 Ordering Provider: Beckman uses the physician name fields. LAW only uses <ID Number>. Beckman and bioMerieux will check what is needed and suggest changes. Recommendation is to add Physician Name Field. This will be enhanced interface data.

- OBR-18 Placer Field 1: Beckman specific data. Beckman started with existing interface and found no place to map requisition field, so used this field. Received from LIS, so use it in reply. May no longer be needed with AWOS ID. Beckman to confirm it is still needed. Assumption is no longer required due to use of AWOS ID.

- OBR-20 Filler Field 1: Beckman specific data. Manual edit of Specimen ID or tube position ID. Made recommendation to indicate original AWOS is an exception, and send result for unsolicited AWOS Status Update with actual Specimen or tube position ID. Beckman will review and consult with IHE LAB. Should not need. No NTE would be needed in the unsolicited result. Recommend not adding this field.

- OBR-21 Filler Field 2: Beckman specific data. Quality Control information. Information needed to fully process control information. Maybe split into OBX and INV. Beckman will determine if INV could be used. Recommend using INV if possible and not adding this field.

- OBR-22 Results Rpt/Status Chng – Date/Time: Release date/time. Could be the same as OBX-19. Beckman will review use of time stamps in OBR, ORC, or OBX to determine if this field is needed. Recommend not adding if existing fields are equivalent.

- OBR-25 Result Status: Release status. Beckman will review and determine if ORC-5 and OBX-11 suffice. Recommend not adding this field if information provided by ORC-5 and OBX-11.

- OBR-32 Principal Result Interpreter: Operator ID of reviewer. Recommend reviewing use of OBX-16 and not adding this field.

- OBR-46 Placer Supplemental Service Information: Number of slides created. Siemens recommended TCD extensions (Siemens extension TCD-10, or other new field in TCD) be used to specify number of runs. Slides have been numbered. TQ1- Quantity is a possibility. Need to add ability to send replicates. May be unique to an instrument, so evaluate this as well. Ultimately may need to propose a change to HL7 OO. Siemens has other extensions for TCD that should be reviewed.