| HL7 Version 2.9 Chapter 13 |  |
| --- | --- |
| Change Request ID: |  |
| File Name: | OO-CR795-TCD-for-LAW.doc |
| Description: | Introduce new fields in the TCD segment required for the IHE LAW profile |
| Status: | New proposal |
| Sponsoring Person |  |
| Sponsoring Business Unit |  |
| Date Originated: | March 25, 2014 |
| Date HL7 approved: |  |
| Backward Compatible: | Yes |
| Forward Compatible: | No |
| HL7 Status & Date |  |

1 Justification Detail:

It is proposed to add three fields to the TCD segment; each of them is described in an own sub-section below.

1.1	Initial Run Count (TCD-10, data type NM)

The LIS may specify the initial count of runs to perform in the given test.

The instrument may decide to perform additional runs as well.

1.2	Auto-Dilution Protocol (TCD-11, data type CWE)

Instead of using a numeric value in TCD-2, the LIS may specify the code of a dilution protocol pre-configured on the instrument, e.g. code “D1” may correspond to dilution 1:20, “D2” — 1:30, etc.

This value is mutually exclusive with TCD-2 “Auto-Dilution Factor”.

1.3	Pool Size (TCD-12, data type NM)

This field is to support the following interaction scenario between an LIS and an instrument:

Test order submission (e.g. OML^O33):

- The LIS sends test orders for patient specimens (SPM-11==”P”).

- Thereby, the LIS wants to specify, that a given specimen may be pooled with other specimens, if circumstances (test type, specimen expiration times, etc.) allow.  For that, the LIS suggests the maximal pool size.

- The instrument understands this parameter, and has a technical possibility to consider it (an own pooler).

Result upload (e.g. OUL^R22):

- Result messages sent by the instrument relate to original specimens, not to the pools (SPM-11==”P”).

- The instrument wants to inform the LIS, that the test result for a given specimen was actually observed when this specimen was pooled with other specimens.  The instrument expresses this by sending the actual pool size, if any.

The proposed new TCD field shall contain the suggested pool size when sent from the LIS to the instrument, and the actual pool size when sent from the instrument to the LIS.

2	Open IssueS:

None known.

3	Change Request Impact:

None known.

4	Documentation Changes:

HL7 Attribute Table – TCD – Test Code Detail

| SEQ | LEN | C.LEN | DT | OPT | RP/# | TBL# | ITEM # | ELEMENT NAME |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  | CWE | R |  | 9999 | 00238 | Universal Service Identifier |
| 2 |  |  | SN | O |  |  | 01420 | Auto-Dilution Factor |
| 3 |  |  | SN | O |  |  | 01421 | Rerun Dilution Factor |
| 4 |  |  | SN | O |  |  | 01422 | Pre-Dilution Factor |
| 5 |  |  | SN | O |  |  | 01413 | Endogenous Content of Pre-Dilution Diluent |
| 6 | 1..1 |  | ID | O |  | 0136 | 01416 | Automatic Repeat Allowed |
| 7 | 1..1 |  | ID | O |  | 0136 | 01424 | Reflex Allowed |
| 8 |  |  | CWE | O |  | 0389 | 01425 | Analyte Repeat Status |
| 9 |  |  | CQ | O |  |  | xxxxx | Specimen Consumption Quantity |
| 10 |  |  | NM | O |  |  | xxxxx1 | Initial Run Count |
| 11 |  |  | CWE | C |  |  | xxxxx2 | Auto-Dilution Protocol |
| 12 |  |  | NM | O |  |  | xxxxx3 | Pool Size |

13.4.10.10	TCD-10 Initial Run Count (NM) xxxxx1 { XE "initial run count" }

Definition: This field contains the initial number of runs the instrument shall perform in the context of the given test order.  Only integer values greater than zero are allowed.  Missing value in this field shall be interpreted as “1”.

The instrument may decide to perform more runs than the initial count specified in this field. When reporting results of each run, OBX-4 “Observation Sub-ID” shall be used to distinguish between them.

13.4.10.11	TCD-11 Auto-Dilution Protocol (CWE) xxxxx2 { XE "auto-dilution protocol" }

Definition: This is an alternative to TCD-2 “Auto-Dilution Factor” for the cases, where an auto-dilution protocol (auto-dilution factor) is pre-configured on the instrument and has a vendor-defined code assigned to it.

The condition for TCD-11 “Auto-Dilution Protocol” is that TCD-2 “Auto-Dilution Factor” is not populated.  It means that these fields cannot be used simultaneously.

13.4.10.12	TCD-12 Pool Size (NM) xxxxx3 { XE "pool size" }

Definition: In order submission messages, this field defines the maximal size of the pool where the given specimen may be combined into.  In result submission messages, this field defines the actual size of such pool.