# Error Handling

# ERR-2 Cardinality

ERR-2: The Cardinality should be [0..*] in table W.3.1-1 since the usage is RE.

# Error Handling Scenarios

In W.2.9, a case for AWOS Request with inconsistent specimen or container information is needed. For example, sending a request for a Specimen Type not supported.  Or, identifying a sample container with carrier information, and identification through a carrier is not supported by the analyzer. This could also be considered a Malformed Message (table values not supported, segment fields not supported) that is handled by an AE acknowledgement. Need to clarify.

In W.2.5.5 Retransmitting Results discuss the SAC, SPM, and Patient segments. If Patient doesn’t match with Order Patient – may want to tie to discussion in W.2.9.6 Inconsistent AWOS Status Update.  Section 3.R.5.3 notes that an Analyzer Manager may retransmit an AWOS request. For example, the AWOS request is sent again, but the patient id is different. This should be an error. NOTE: Agreement was reached on this approach. AWOS is tied to patient, SPM, container, and test (PID, SPM, SAC, OBR).

# Coding Systems

In some cases cannot APRIOR identify the entire set of codes. Should this be differentiated? Review OBX-3 versus SPM-11. Service IDs are not completed specified, but Specimen Roles are.

# Length of ERR Text Fields

When CE and CWE is used, HL7 2.5.1 defines length of ST data type as 199. Need to indicate this is truncatable.

# Support for additional ERR fields

LAW should support the following fields in order to provide guidance on application errors in addition to malformed message errors:

- ERR-5 Application Error Code, usage RE, (requested by Beckman Coulter)

- ERR-7 Diagnostic Information, usage RE

- ERR-8^ User Message, usage RE, (requested by Beckman Coulter)

“RE” should be used because not all systems may be able to respond with this information.