# Standardized Graphs

# Scatterplots and Histograms Background

- Very few Analyzer Managers have implemented reconstruction from points

- No customers have asked for access to the points

- Analyzer Manager wants the image

  - Eliminate need to reconstruct

  - Easier for other vendors

- Manual diffs are always needed (5 to 20% of tests)

- People performing reviews doubled efficiency with access to electronic image

  - e.g. from 4 per hour to 8 per hour

  - No paper (printing of plots)

  - No filing of paper, recoded in instruments

- Prior information important as well, because can look at the previous electronic result

- Standardize on the image

# Standardized Image for Graphs

- Support small set of codes

- 800x600, 5k of data, thumb nail resolution

- May be part of the image meta data

- Image Size relates to size of the file to transmitted

- Any reason to standard on size?

  - Recommended 800x600

- Can link to a more dense image

  - use RP

  - Clarify how Data types can be used for Supplement Results

- Limited or no meta data

  - Title (The OB-3.Text field?)

    - Text

  - Category (histogram, scatterplot, cell type)

    - Sortable

  - Associated information (notes, comments)

    - Text

- Use multiple OBX, or part of metadata

- Cellavision sends a 600x600, 500K, very detailed

- Some amount of availability of higher resolution file

- OBX-5

  - Use ED datatype

  - Identify appropriate image data types

    - JPG

    - PNG

    - BMP

- Review length restrictions on OBX, update to support images

- OBX-8, OBX-11

  - Plot captures multiple observation

  - What should be populated?

  - What does an F mean for an S_IMAGE? The image is complete – not the results

  - Highlight NULL can be sent for OBX-8

- Should Reference Pointer also be an S_????

- Beckman usage of fields for Images

  - OBX-3.1 ID

  - OBX-3.2 Human Readable

  - NTE that follows – Associated information

    - Eliminates “pressure” on length of OBX

    - Not as clean due to segment dependency

# Points for a graph

- Scatter plots have large number of points

- Histograms have small number of points

- Tightly coupled with instrument

- Forces additional requirements on the Analyzer Manager

- Recipient must reconstruct

- Analyzer Manager may be able to provide additional capabilities with individual points

- Some common graph definition/recommendations may be possible (XML, JSON)

- Difficult to standardize across instruments

- Will not be standardized by LAW

- Discusses Compression?