TCD-in-LAW call on 25 March 2014 at 5 pm CET

http://wiki.ihe.net/index.php?title=Laboratory_Analytical_Workflow#OBR_Fields

Attendees: Joanna Selinsky, Edwin Heierman, Dmytro Rud

Updated documents:
- Proposal for HL7: http://wiki.ihe.net/index.php?title=File:795.docx
- Proposal for IHE: http://wiki.ihe.net/index.php?title=File:CP-LAB-zzz-LAW-TCD.docx


TCD-9 “Result Aspect”
Clarification from Siemens (Dilip Jha) received:
=======
Q: What is the intention of this field? Is this an attribute of a test order submission, of an observation result, or of both? Which actor is supposed to send it in which cases?

Dilip: One result produced by instrument can have multiple interpretations like qualitative (like HIV +ve or -ve) and quantitative (numerical score like 2.5). In this case one result has two result aspects. Since it is result, it is part of Observation result (OBX) only. Instrument is supposed to send this to LIS Manager. In this case two results need to be sent to LIS Manager with different result aspects (one with say I call “Quantitative” other “Qualitative” for simplicity) and both are final result.

Q: If this is an attribute of a test order submission: would it be possible to use OBR-46 instead? 

Dilip: No

Q: If this is an attribute of an observation result: can it be considered a part of supplemental results as described in section W.2.5.4 of the LAW specification?

Dilip: not a supplemental result, both are final result.
=======

Further discussions required.


TCD-10 “Replicate Number”
Decision: use TQ1-2 “Quantity” instead, with the optionality RE on Analyzer Manager and O on Analyzer.

[bookmark: _GoBack]Pending action (Dmytro): update LAW section X.2.4 “Rerun” ???

Open question: Is an additional Use Case in the LAW profile required?  
To be discussed with François.


TCD-11 “Dilution Protocol”
Decision: accept as Conditional (RE/X) on Analyzer Manager and Not Supported on Analyzer, rename to “Auto-Dilution Type”, data type CWE, vendor-specific values expected.

Open questions: Does the dilution type/protocol denote only the dilution factor or other dilution modalities as well? Is it really mutually exclusive with TCD-2 “Auto-Dilution Factor”?


TCD-12 “Pool Size”
Decision: accept as RE on Analyzer Manager and O on Analyzer, move to TCD-10.

Open question: Is an additional Use Case in the LAW profile required? 
To be discussed with François.


